# Sensor de ph
Módulo ph4502-c

documentación y bibliotecas
https://github.com/nthnn/PH4502C-Sensor

conexiones a arduino uno
https://www.taloselectronics.com/blogs/tutoriales/como-usar-modulo-ph-4502-con-arduino

video interesante
https://www.google.com/search?q=ph4502c+arduino+desde+cero&client=ubuntu-sn&hs=9Rp&sca_esv=04c57bc3f34f5e13&channel=fs&sxsrf=ADLYWIKIxsb1zBn4G0lnQpXkeayCnhxR-w%3A1717492637369&ei=ndteZs2aFtzT1sQPlvCIoA4&ved=0ahUKEwiN3bfUzsGGAxXcqZUCHRY4AuQQ4dUDCBA&uact=5&oq=ph4502c+arduino+desde+cero&gs_lp=Egxnd3Mtd2l6LXNlcnAiGnBoNDUwMmMgYXJkdWlubyBkZXNkZSBjZXJvMgUQIRigATIFECEYoAFImH1Qog5YmXZwBHgBkAEAmAF5oAH9E6oBBDI4LjG4AQPIAQD4AQGYAiGgArYUwgIEEAAYR8ICChAAGLADGNYEGEfCAgQQIxgnwgIIEAAYgAQYogTCAgoQIxiABBgnGIoFwgIFEAAYgATCAgQQABgewgIKEAAYgAQYFBiHAsICBxAAGIAEGBPCAgcQABiABBgNwgIGEAAYFhgewgIHECEYoAEYCsICBBAhGArCAgQQIRgVmAMAiAYBkAYIkgcEMzAuM6AH9FY&sclient=gws-wiz-serp#fpstate=ive&vld=cid:08596d5d,vid:XqudFPx29SM,st:0
# PH4502C Sensor Arduino Library
This Arduino library provides a convenient and easy-to-use interface for the PH4502C sensor, allowing you to effortlessly integrate it into your Arduino projects. The PH4502C sensor is a high-precision pH sensor capable of accurately measuring pH levels in various applications.

## Features
- Simple and intuitive API for interacting with the PH4502 sensor.
- Supports pH measurement with high precision and accuracy. Automatic temperature compensation for more accurate readings.
- Flexible calibration options to ensure optimal performance.
- Lightweight and efficient implementation for resource-constrained Arduino boards.
## Installation
To use the PH4502 Sensor Arduino Library, follow these steps:
### Installation on Arduino IDE
Just type "PH4502C-Sensor" on the Library Manager of the Arduino IDE.

Alternatively you can follow the steps below:
Download the library by clicking on the "Download" button or clone it using Git:
git clone https://github.com/nthnn/PH4502C-Sensor.git
Place the downloaded library folder into the Arduino libraries directory on your computer. In Windows, the default location is usually:
[Arduino Documents Folder]/libraries/
Launch the Arduino IDE. Then, click on "Sketch" and then navigate to Include Library. You should see the newly added "PH4502-Sensor" option. Click on it to include the library in your sketch.
Now you are ready to use the PH4502 Sensor Library in your Arduino projects.
### Installation on PlatformIO
Type "PH4502C-Sensor" on the PlatformIO library registry, and click "Add Project." And that's it!

## Getting Started
To begin using the PH4502 Sensor Library, follow these simple steps:
1. Connect your PH4502C sensor to your Arduino board. Make sure to connect the appropriate VCC, GND, and data pins based on your Arduino model and pin configuration.
2. In your Arduino sketch, include the PH4502C Sensor Library by adding the following line at the top of your code:
  #include <ph4502c_sensor.h>
Create an instance of the PH4502C_Sensor class:
PH4502Sensor ph4502c(A0, A1);

// Analog Pin 0 - pH level pin
// Analog Pin 1 - temperature pin
Initialize the sensor in the setup() function of your sketch:
void setup() {
    // Initialize the sensor
    ph4502c.init();
}
Read the pH value from the sensor using the readpH() function:
void loop() {
    // Read pH value
    float pHValue = ph4502c.read_ph_level();

    // Do something with the pH value
    // Delay before next reading
    delay(1000);
}
Optionally, you can calibrate the sensor for better accuracy. Refer to the library documentation for more information on calibration.

# Como usar modulo PH-4502 con arduino?
El módulo PH-4502 es una herramienta esencial para aquellos que requieren mediciones precisas y confiables de pH en sus proyectos electrónicos.
Especificaciones
- Alimentación en 5Vdc
- Consumo de corriente: 5 a 10mA
- Temperatura de trabajo: -10 a 50 °C
- Conexión de sensores estándar mediante conector BNC
- Dimensiones: 42 x 32 x 20 mm
- Peso: 25g
Características
Amplio Rango de Medición: Con un rango de medición típico de 0 a 14 pH, este módulo es apto para una amplia variedad de aplicaciones que abarcan desde soluciones altamente ácidas hasta soluciones altamente alcalinas.
Interfaz Sencilla: Su interfaz de conexión sencilla permite una integración rápida y eficiente con plataformas electrónicas como Arduino, facilitando la adquisición y procesamiento de datos de pH.
Salida Analógica y Digital: El módulo proporciona tanto una salida analógica como una salida digital, lo que brinda flexibilidad en la forma de utilizar los datos obtenidos.
Calibración Sencilla: El módulo PH-4502 es compatible con procedimientos de calibración estándar, lo que garantiza mediciones precisas y confiables tras la calibración inicial.
Amplio Espectro de Aplicaciones: Este módulo es ideal para aplicaciones que requieren monitoreo y control de pH, como acuaponía, sistemas de riego automatizado, acuarios, laboratorios, entre otros.
Material
- Arduino uno
- Modulo PH-4502
- Sonda PH

El modulo PH-4502C Sensor de PH es un dispositivo que permite medir el PH con ayuda de una sonda que es la que toma la lectura (electrodo E201) mediante el conector BCN.
Cuenta con un potenciómetro de offset el cual te ayuda a calibrar la medición de PH (el mas cercano al conector BNC) mientras que el otro potenciómetro controla la salida Do la cual se activa cuando el PH haya alcanzado el limite que hayamos configurado a través del potenciómetro cabe mencionar que la salida Do es de 3.7V como salida alta.
El pin Po puede conectarse directamente a un pin analógico de un Arduino, PIC o cualquier tecnología TTL, por lo que no es necesario alguna librería especifica para que podamos utilizarlo. Como adicional cuenta con dos indicadores LED uno para señalar la alimentación del circuito y otro para señalar la salida Do.
Conexiones
- VCC del sensor -> VCC del módulo
- GND del sensor -> GND del módulo
- Salida analógica del sensor -> A0 del módulo

Calibración
Para calibrar el sensor haremos un procedimiento muy sencillo, cortocircuitando el pin central del BNC con la parte externa del conector, lo que hará que cuando conectemos el módulo al Arduino nos mostrará el valor medio que se convirtió en voltaje. debe ser 2,5V.

Codigo
/*
* Calibracion modulo PH
*/

int entrada_A0;
float voltaje;

void setup() {
  Serial.begin(9600);
}

void loop() {
  entrada_A0 = analogRead(A0);
  voltaje = (entrada_A0 * 5.0) / 1024.0;
  Serial.println(voltaje);
  delay(500);
}

Cargaremos el código y abrimos el monitor serie, Este nos tiene que dar un valor cercano a 2.5 si no lo esta dando entonces tendremos que mover el potenciometro hasta llegar al voltaje.
Hasta aqui se temina este tutorial, mas adelante integraremos una sonda y mediremos el valor de ph.

# How to use a PH probe and sensor
If you worked with PH metering before you will know that PH values range from 0-14. Where
PH 0 Will be very acidic, PH 7 will be neutral and PH 14 very alkaline. Water is near a PH 7 and
this is usually around here that we will need to monitor PH of many things. A swimming pool,
for example, should be slightly alkaline at 7.2, hydroponics systems around 6 (for optimum
plant nutrition takeup) and aquaponics around 6.8.
I wrote this PH probe and sensor “how to” because it is not as straightforward as one would
think (but quite easy when you understand the ins and outs) mostly because there is not a
lot of information on this on the Internet, surely not detailed information.
We will first look at the ph probe module board and then the PH probe because both the PH
probe and sensor have to be set correctly:
 offset setting
 limit setting
 sketch to test the board analogue range
 sketch for PH reading and calibration.
 calibration of PH probe
 PH probe usage
The ph probe module in this tutorial is available on our site here: PH probe module BNC
conector

PH Probe Sensor Pinout
TO – Temperature output
DO – 3.3V Output (from ph limit pot)
PO – PH analog output ==> Arduino A0
Gnd – Gnd for PH probe (can come from Arduino GND pin) ==> Arduino GND
Gnd – Gnd for board (can also come from Arduino GND pin) ==> Arduino GND
VCC – 5V DC (can come from Arduino 5V pin) ==> Arduino 5V pin
POT 1 – Analog reading offset (Nearest to BNC connector)
POT 2 – PH limit setting
PH probe module Offset and how to use it.
This board have the ability to supply a voltage output to the analogue board that will represent
a PH value just like any other sensor that will connect to an analog pin. Ideally, we want a PH
0 represent 0v and a PH of 14 to represent 5V.
BUT there is a catch……, this board by default have PH 7 set to 0V (or near it, it differs from
one PH probe to another, that is why we have to calibrate the probe as you will see later on),
This means that the voltage will go into the minuses when reading acidic PH values and that
cannot be read by the analog Arduino port. The offset pot is used to change this so that a PH 7
will read the expected 2.5V to the Arduino analog pin, the analog pin can read voltages
between 0V and 5V hence the 2.5V that is halfway between 0V and 5V as a PH 7 is halfway
between PH 0 and PH 14,
You will need to turn the offset potentiometer to get the right offset, The offset pot is the blue
pot nearest to the BNC connector.
To set the offset is easy. First, you need to disconnect the probe from the circuit and short-
circuit the inside of the BNC connector with the outside to simulate a neutral PH (PH7). I took
a piece of wire, strip both sides, wrap the one side around the outside of the BNC connector
and push the other side into the BNC hole. This short-circuit represents about a neutral PH
reading of 7.

There are two ways you can do the adjustment.
If you have a multimeter handy you can measure the value of the PO pin and adjust the offset
potentiometer until PO measures 2.5V.
I prefer to just use the sketch below. Just download it to your Arduino as you will with any
other sketch, open serial monitor and view the reading there. All this sketch does is to print
the volts it receives from the analog pin and print it to the serial monitor. It of course first
changes the digital value to volts to make it easier. Now simply turn the offset pot until it is
exactly 2.5V. You can learn more about reading voltages and digital representation of volts
here: https://www.arduino.cc/en/Tutorial/ReadAnalogVoltage
Offset sketch
void setup() {
// initialize serial communication at 9600 bits per second:
Serial.begin(9600);
}
// the loop routine runs over and over showing the voltage on A0
void loop() {
// read the input on analog pin 0:
int sensorValue = analogRead(A0);
// Convert the analog reading (which goes from 0 - 1023) to a voltage (0 - 5V):
float voltage = sensorValue * (5.0 / 1023.0);
// print out the value you read:
Serial.println(voltage);
delay(300);
}
PH limit setting
There is another pot that acts like a limit switch. Basically, the D0 pin on the sensor board will
supply 3.3V to the pin until a preset PH value (that you set with the limit pot) is reached, at
this point a red LED will light up and the pin will go down to about 0V.
I did not play with this much but suppose it can be handy if you want to activate a buzzer or
something if a certain PH is reached, it will work great on an Arduino digital port – that will go
high from about 2V up.
This will work if the PH value goes higher than the set value. If you want it to trigger
something when the PH goes lower, you need to monitor the digital pin to trigger when the
digital pin goes low.
You will unfortunately not be able to set this limit between two values, either if the pH goes
up to high or if the PH drop to low. Programmatically you of cause can do an upper and lower
limit.
Connecting and calibrating the PH probe.
The hard part is over and this offset does not have to be set again, even if you change PH
probes. We have PH probes available here: PH probe Electrode BNC connector
Here is a couple of things to know about PH probes:
1. The probes readings change over time and need to be calibrated every now and again to make
sure the value is still the same and be adjusted if it did change.
2. You need at least one PH buffer solution to calibration your PH probe. They are available at
many different PH values, A buffer solution of 6.86 and 4.01 is most common as it covers the
range of most applications. If you are only going to use one buffer solution make sure its value
is near the value range you will use in your normal tests – if it is pool water a buffer solution of
6.86 is usually near enough.
3. Buffers come in pre-made solutions or as a powder. I prefer the powder because it is cheaper
and does not have an expiration date. The powder is easy to make up as well, I suppose it
depends on the power you will use, the one I use you add the powder to 250ml distilled water
and stir until all powder is dissolved. It will last about a month once you added water to it.
A PH probe takes some time to get to the right value, allow it to be in the liquid you want to
measure for at least two minutes or longer, it does not mean it will be stable at one ph value,
it will jump around a bit between 3 or 4 values but on the last digit, for example, between
6.84 – 6.88
4. PH values differ in different temperatures, although that might sound cumbersome, in the
temperature range between 10 – 30 degrees Celcius the PH does not differ and from 30
degrees Celcius it goes up with about a pH of 0.01 to 50 degrees Celcius that is about 0.06. In
most uses, it will be below 30 degrees Celcius and temperature do not have to be calculated
in.
Hook up your PH probe after you removed the wire you used to short-circuit the BNC
connector and download the sketch below.
PH measurement sketch
float calibration = 0.00; //change this value to calibrate
const int analogInPin = A0;
int sensorValue = 0;
unsigned long int avgValue;
float b;
int buf[10],temp;
void setup() {
Serial.begin(9600);
}
void loop() {
for(int i=0;i<10;i++)
{
buf[i]=analogRead(analogInPin);
delay(30);
}
for(int i=0;i<9;i++)
{
for(int j=i+1;j<10;j++)
{
if(buf[i]>buf[j])
{
temp=buf[i];
buf[i]=buf[j];
buf[j]=temp;
}
}
}
avgValue=0;
for(int i=2;i<8;i++)
avgValue+=buf[i];
float pHVol=(float)avgValue*5.0/1024/6;
float phValue = -5.70 * pHVol + calibration;
Serial.print("sensor = ");
Serial.println(phValue);
delay(500);
}
A note on buffer solutions: do not CROSS CONTAMINATE! What I mean by this is to not take
the probe from one buffer solution to another or from a liquid sample you tested to a buffer
solution without rinsing it thoroughly with distilled water first. You will change your buffer
solutions ph and your calibration will be off.
When you place the probe in the first solution you might be surprised at how far off it can be,
it’s normal though. Remember to leave the probe in the solution for at least two
minutes to stabilise. In the script’s top line you will see a variable
called “calibration”. Change this value to the difference between what you see in serial
monitor and the buffer solutions value, for example, if you read 5.81 and the buffer solution is
6.86 you should change the variable’s value to 2.05.
Upload the changed sketch and see how the value looks now.
Some ideas for you.
You can add a potentiometer to your project and program that to change the calibration for
you. You always run a risk that the pot might be adjusted my mistake so a button with a 5-
second delay can be programmed to put the unit in calibration mode.
This becomes great if you add an LCD screen to it and if you add the calibration value in
EEPROM so it holds it even if the PH project is powered off.
How about adding a buzzer to the 3.3V limit output to notify you when the PH is out of range.
Usually an upper or lower will be enough, in most applications, you will either be worried about
high or low PH values but not both. If this is a problem and you need the PH to be in a certain
range you can easily do that programmatically and have the buzzer on a digital pin. If it is
crucial you can even add a GSM module to this so you can get an SMS or how about a dosing
pump to automatically get the PH back in range.
I just might, but do not hold your breath (so much to do but so little time) :-), do something
like that in the future with our cheap Skeleton Duino.
PH probe and sensor conclusion
If you go through these steps once it becomes as easy as pie. All the parts, even those I
mention in the ideas section is available on our site.

# Experimentos
los datos se reistran en data.frame: datos
columnas
> names(datos)
"fecha"	fecha del experimento
"electrodo"	electrodo utilizado
"segundosmedicion" segundos que se tomó la medicion
"ph" ph de la solucion testeada
"mvolt" mvolts registrado por arduion
"lote"	lote de experimentacion. Un mismo lote es un experimento
 "orden" orden en que se hicieron las mediciones dentro de un lote
## Lote1 - estabilizacion lectura y memoria
Objetivo: ver tiempo de estabilización e inercia al cambiar de solucion
medición de pH com electrodo viejo (azul). Se midio durante 100 segundos (cada 10s) solucion de pH=10, dos veces consecutivas y luego la de pH=4
Los datos se hallan en data.frame: datos en carpeta phmetro dentro de cuem/arduino/equipamientoc
codigo:
max=max(datos$mvolt[datos$lote==1],na.rm=TRUE)
min=min(datos$mvolt[datos$lote==1],na.rm=TRUE)
plot(datos$segundosmedicion[datos$lote==1 & datos$ph==10 & datos$orden==1],datos$mvolt[datos$lote==1 & datos$ph==10 & datos$orden==1],ylim=c(min,max),pch=19,las=1,ylab="mVolts",xlab="segundos",main="electrodo viejo")
points(datos$segundosmedicion[datos$lote==1 & datos$ph==4],datos$mvolt[datos$lote==1 & datos$ph==4],pch=19, col="red")
legend('topright',legend=c('ph 10','ph 4'),pch=19,col=c('black','red'))

Conclusion: en 60 segundos parece haber una buena estabilización y los valores se diferencian bien de pH=4 a pH10.
## Lote2 - curva calibración
Objetivo: curva de calibracion
se utilizo electrodo (viejo - azul) se midio pHs: 4-6-8-10
tiempo de medicion: 100s
Los datos se hallan en data.frame: datos en carpeta phmetro dentro de cuem/arduino/equipamiento
max=max(datos$mvolt[datos$lote==2],na.rm=TRUE)
min=min(datos$mvolt[datos$lote==2],na.rm=TRUE)
plot(datos$ph[datos$lote==2],datos$mvolt[datos$lote==2],ylim=c(min,max),pch=19,las=1,ylab="mVolts",xlab="segundos",main="electrodo viejo")

lmlote2<-lm(datos$mvolt[datos$lote==2]~datos$ph[datos$lote==2])
summary(lmlote2)

Call:
lm(formula = datos$mvolt[datos$lote == 2] ~ datos$ph[datos$lote ==
    2])

Residuals:
   Min 1Q Median 3Q Max
-2.616 -1.576 -1.147 1.282 5.483

Coefficients:
        Estimate Std. Error t value Pr(>|t|)
(Intercept) 2720.2776 3.2875 827.457 < 2e-16 ***
datos$ph[datos$lote == 2] -3.3502 0.4638 -7.224 0.000357 ***
---
Signif. codes: 0 ‘***’ 0.001 ‘**’ 0.01 ‘*’ 0.05 ‘.’ 0.1 ‘ ’ 1

Residual standard error: 2.84 on 6 degrees of freedom
Multiple R-squared: 0.8969,	Adjusted R-squared: 0.8797
F-statistic: 52.18 on 1 and 6 DF, p-value: 0.0003568

Conclusion: existe correlación entre mvolts y pH con buen ajuste segun el lm()
## Lote 5 - curva calibración
Se monto placa ph4502c con Arduino uno en caja plástica con puertos USB B de arduino y BNC para el electrodo. Se probó con soluciones estándar 4 - 6 - 7 y 10. Al finalizar la curva se repitió la medida de la solución 4.
tiempo de medición 100 segundos. El sketch toma 500 mediciones insatntaneas la promedia y almacena, repite el proceso cada segundo durante los 100 segundos. Por lo tanto cada medicion es el promedio de 50000 mediciones.
Advetencia: los valoes de mVolts comparados con Lote 1 son 10 veces menores debido a que en el sketch se cambió un divisor de 100 a 1000.
Codigo
max=max(datos$mvolt[datos$lote==2],na.rm=TRUE)
max=max(datos$mvolt[datos$lote==5],na.rm=TRUE)
min=min(datos$mvolt[datos$lote==5],na.rm=TRUE)
plot(datos$ph[datos$lote==5 & datos$orden<5],datos$mvolt[datos$lote==5 & datos$orden<5],ylim=c(min,max),pch=19,las=1,ylab="mVolts",xlab="pH",main="electrodo viejo")

modelo lineal
> lmlote5<-lm(datos$mvolt[datos$lote==5 & datos$orden<5]~datos$ph[datos$lote==5 & datos$orden<5])
summary(lmlote5)

Call:
lm(formula = datos$mvolt[datos$lote == 5 & datos$orden < 5] ~
    datos$ph[datos$lote == 5 & datos$orden < 5])

Residuals:
        1 2 3 4
-0.003733 0.070800 -0.086933 0.019867

Coefficients:
        Estimate Std. Error t value Pr(>|t|)
(Intercept) 267.8928 0.1319 2031.39 2.42e-07 ***
datos$ph[datos$lote == 5 & datos$orden < 5] -0.3323 0.0186 -17.86 0.00312 **

(Intercept)
datos$ph[datos$lote == 5 & datos$orden < 5]
---
Signif. codes: 0 ‘***’ 0.001 ‘**’ 0.01 ‘*’ 0.05 ‘.’ 0.1 ‘ ’ 1

Residual standard error: 0.08056 on 2 degrees of freedom
Multiple R-squared: 0.9938,	Adjusted R-squared: 0.9907
F-statistic: 319 on 1 and 2 DF, p-value: 0.00312

Conclusión: El ajuste es muy bueno y el valor de la pendiente repitió muy bien comparado con lote1. Lote 1 pendiente: -3,3502 Lote 5 pendiente= -0,3323. La diferencia de 10 se debe solo a un cambi de un divisor dentro del sketch

Se construyo funcion de mv~ ph y se graficó la recta junto con los puntos
mvlote5<-function(ph){267.8928-0.3323*ph}
max=max(datos$mvolt[datos$lote==5],na.rm=TRUE)
min=min(datos$mvolt[datos$lote==5],na.rm=TRUE)
plot(mvlote5,4,10,las=1,ylab="mvolt",xlab="pH")
points(datos$ph[datos$lote==5 & datos$orden<5],datos$mvolt[datos$lote==5 & datos$orden<5],pch=19)

Conclusion: se ve un buen ajuste.
## Lote3
Objetivo: inercia o memoria de la medicion al cambiar de solucion electrodo viejo-azul
se utilizo electrodo (viejo - azul) se midio pHs: 4-10 y se volvio a 4
tiempo de medicion: 100s
Los datos se hallan en data.frame: datos en carpeta phmetro dentro de cuem/arduino/equipamiento
> min=min(datos$mvolt[datos$lote==3],na.rm=TRUE)
> max=max(datos$mvolt[datos$lote==3],na.rm=TRUE)
> plot(datos$orden[datos$lote==3],datos$mvolt[datos$lote==3],ylim=c(min,max),ylab='mvolts',xlab='orden de medicion')

Conclusión: al pasar de 4 a 10 hay respuestay luego al volver a 4, se repite el valor
## Lote4
Objetivo: curva de calibracion e inecia, electrodo nuevo-negro
se utilizo electrodo (nuevo - negro) se midio pHs: 4-6-8-10 y a continuacion pH 4 nuevamente
tiempo de medicion: 100s
Los datos se hallan en data.frame: datos en carpeta phmetro dentro de cuem/arduino/equipamiento

# end